
OPERATING ROOM TABLE SHIELD TECHNICAL SPECIFICATION

1. The system must be suitable for the table used in the operating room.
2. System eaves must consist of at least two sections.
3. For mounting the system on the table, there must be a special apparatus that is specially coated and will not cause any reactions in the form of rust or mold.
4. After the system is mounted on the table, it must have the ability to move left, right, forward and backward. The same mechanism should be suitable for use on the right or left of the table, if desired.
5. The width of the system should be 120 cm +-5 and the height of the system should be 60 cm +-5.
6. The minimum protection ratio must be equivalent to 0.50 mm Pb.
7. The system mounted on the table will not obstruct the movements of the table.
8. The system mounted on the table will not obstruct the movements of the table.
9. The x-ray protective lead curtain must be produced and certified in accordance with the Personal Protective Equipment Regulation (Official Gazette No. 30761 dated 01 May 2019, new Regulation (EU) 2016/425
10. In line with the old standards for the products, Articel 11 A, Articel 11 B certificates should not be submitted. These documents have become obsolete and documents must be submitted for the sale of products certified according to the new revision.
11. In line with the old standards for the products, Articel 11 A, Articel 11 B certificates should not be submitted. These documents have become obsolete and documents must be submitted for the sale of products certified according to the new revision.
12. There must be TAEK approval and test reports for the production of the product. It should offer this when requested. A report showing at least 96% protection of the product must be submitted with the product.
13. The production conditions for the product must comply with ISO 9001:2015 standards and a document must be submitted proving that it complies with these standards.



